## Slide 1
MinION Anatomy

## Slide 2
Pre-Run Check List
Flowcell inspection Remove bubbles in fluid lines and on surface of flowcell where possible if present, & confirm conductive heat pad is installed on bottom surface of ASIC chip.
4
Check computer SSD for available storage space >150Gb
3
Make sure required Software is installed:-
MinKNOW Control of MinION device & run parameters
Metrichor Cloud basecalling of event data
MinoTour Live monitoring / control of run while sequencing
Chronolapse Screen image grabber for record keeping
Poretools, poRe Sequence extraction and data summaries
TeamViewer Remote control of MinION computer
1
Confirm automatic software updates and sleep modes are disabled
2

## Slide 3
Attach Flowcell to MinION device and plug into USB3 port
Start MinKNOW software and start device
Name Run and start platform-QC protocol
If flow-cell good proceed
Prime flow-cell with 2x 500ml 1xRBF1 for 10 mins each
(500ml 2xRBF1 + 500ml H2O)
Load 150ml of Library in 1xRBF1
(75ml 2xRBF1 + xml H2O + 4ml Fuel + 6-11ml Pre-sequencing Mix)
Name run and start Sequencing Protocol – Standard / Modified :o)
(Start Metrichor & required workflow plus screen capture software)
Preparing & Loading MinION Device
1
2
3
4
5
6
(Single Good pores >650)

## Slide 4
Inserting Flow-cell into MinION Device
Ready MinION
Open lid
Unpack flow-cell, check heat pad
intact & ASIC bubble free
Slide flow-cell into place
Make sure flow-cell is
seated properly
Plug MinION into USB3 port
1
2
3
4
5
6

## Slide 5
Flow-cell Quality Assessment
Good
Bad
Platform-QC scans the 2048 channels as 4 Mux groups @ -180mV with R9 & reports back single good channel assignment to each of g1 to g4 groups with up to 512 wells each. Expect Good correlation using R9 flowcells between Platform QC pore numbers and those obtained at the start of a run.
>650 Pore Guarantee

## Slide 6
Priming and Loading Flow-cell
Open sample port
Slowly remove any air from sample port with pipette
Confirm fluidic channel is free of bubbles
2 x 10 min 500ml 1xRBF1 followed by 150ml Library
Close sample port
Close lid and GO!
1
2
3
4
5
6

## Slide 7
Priming and a Loading SpotON Flow-cell
Open sample port
Slowly remove any air from sample port with pipette
Add 75ul of Sequencing Mix to SpotON well dropwise
Open SpotON port and then add 200ml RBF1 through sample port, NOT SpotON port.
Close ports, lid
and GO!
1
2
3
4
1x 10 min 500ml RBF1 through sample port
1x 10 min 300ml RBF1 through sample port
Flowcell Priming

## Slide 8
Start Run
Start Sequencing Recipe Script
Name run and select required Recipe Script
Start Chronolapse screen capture
Image grab every 30 seconds and as required at start
Start Metrichor workflow
Select desired workflow:- 2D basecall or WIMP
1
2
3
Start MinUp & Open MinoTour Browser Window
Issue command line parameters for read mapping and device control as below or use MinUP-GUI:
minUP.exe -dbh minotour.nottingham.ac.uk -dbu USER -pw PWD -w d:\data -f c:\reference\REF.txt -u USER -c -bwa -d -ip xxx.xxx.xxx.xxx -pin XXX -s minion
4

## Slide 9
Good or Bad Library/Flowcell ?
95%
90%
97%
96%
91%
96%
84%
97%
93%
97%
97%
17%
83%
80%
83%
55%
92%
92%
86%
80%
% of Single Good pores in Strand

## Slide 10
Single Good Pore % in Strand as Library QC
292 / 302 = 97%
Use the relative quantity of in
Strand reads as a measure of
library quality independent
of pore number

## Slide 11
Script Tinkering
Bias-Voltage Setting & Remux
Bias voltage directly controls induced current flow across the membrane, and current flow is used to assign pores to different categories. With time greater bias-voltage is required to produce current that is in the “single good” pore range. Bias-voltage is the master control. Selection or “re-mux” of pores at a particular targeted voltage must be carried out for maximum efficiency.
1
Yield Monitoring
As pore numbers fall and/or reagent is depleted you will see a drop in the event yield over time. Gaining access to still functional, but unselected, pores once event accumulation rate drops below a set value makes much better use of a flow-cell. This lessens unproductive electrochemical gradient deterioration observed with standard recipe scripts.
2
Pore Shepherding
Wells actively sequencing experience an electrochemical gradient deterioration that is twice that of inactive pores, but even inactive pores show a deterioration that requires bias-voltage increases for optimal functioning. Because of the 4-1 (wells - electrical channel) nature of current flow-cells this results in a spreading of the optimal bias-voltage required for a particular well depending on its on/off history. To mitigate this, sub-peak bias-voltage selection can “shepherd” well population more tightly by attempting to deplete least used wells.
3

## Slide 12
Standard 48hr Sequencing Recipe Script
g1 to g2 switch
24hrs
g1
8hrs
R9 48hr Genomic DNA Sequencing Script
Mux @-180mV to assign g1, g2, g3 & g4 pores
Example Pore Usage from retired R7.3
g2
8hrs
g3
8hrs
g4
8hrs
g1
8hrs
g2
8hrs
STOP
(example from previous R7.3 Version)
…

## Slide 13
Modified Scripts to Maximise Flowcell Yields
…..
Yield dependent
time
After remux voltage scan
After remux voltage scan
After remux voltage scan
Yield dependent time
Yield dependent time
Yield dependent time
Switch when hourly
event yield drops
below 67% of first
hour of this segment
Multi-Voltage Re-mux
scan selecting one
within a defined
range of the greatest
“single good pore”
count voltage
Yield-Monitoring_Voltage-Tuning Scripts
(Voltage tuned across 30mV
range)
Mux / Re-Mux Scan
15 min
Yield Dependent Time
x5 (user defined)
(~1-5 hrs)
-180mV
Start with fresh flow-cells
(If new flowcell use -180mV to start)
Example Pore Usage

## Slide 14
Run Dynamics:- SQK006 Standard Recipe Script
Event Accumulation
Abasic Current
Pore Usage

## Slide 15
Run Dynamics:- SQK006 Tuning Recipe Script
Abasic Current
Pore Usage
Event Accumulation

## Slide 16
Standard Script
(Chip65)
Tuning Script
(Chip66)
Individual Pore Read Starts
Individual Pore Read Starts